Alicat Firmware Version 6+

Data Frame Query Commands

With the upgrade to firmware version 6v0 and the inclusion of user selectable units of measure, the output of the data frame query command, (unitID)??D*, was altered to accommodate the different ways that the data can be presented. There were three main factors which made retaining the original output of this data frame query a non-trivial task: the addition of new possible variables in the data frame, variables showing up in customized order on the data frame, and the units of measure being fully selectable. As such, the information in the data frame query was altered with the change to version 6 of the firmware.

If the ??D* query needs to remain unchanged, please let us know and we can put firmware version 5v12 on the device (losing the selectable units of measure, but retaining the original data frame query output).

The original data returned from the data frame query command was in the following format:

The new format for the data returned from the data frame query command is as follows:

Each variable now has a unique ID number associated with it internally, which can be used to associate the data frame variable with the appropriate measurement type based on the ID numbers. The min and max values were removed from this output due to the minimum values not necessarily being accurate due to how information was stored relating to the calibrated span of the device. The max value was removed due to it now having the possibility of containing many more digits than would fit in the MaxVal column as well as being dependent on the units of measure you are referencing. The units column was expanded to include more potential information, and the specific units of measure now also have a unit ID associated with them (tied to the variable with which the units are associated with).

The software version can be queried using a command of (unitID)VE. The response to this will contain the main software version, the subversion, the board version (R22 being a revision 22 main circuit board in the image below), and the date of the software build. Below is an example:

Please see the following pages for a list of the variable ID’s and their associated engineering units ID’s (for software version 6v0 and higher).

For quickly determining the variable and units of measure in software version 6v0 and higher, there is now also a faster truncated data frame query command, (unitID)??D+, whose outputs are formatted as below:

In this truncated version, the only information provided is the ID of the variable/status that is expected to show up (statuses will only show up if they are triggered, such as the MOV error) and the units of measure if it is a data variable. The units of measure here are indicated by the number ID associated with the units list of the particular variable (see following pages).

Because the full scale value is now dependent on the units of measure and can be larger than previously expected, there is a new command designed to specifically read out the full scale of a given variable (in either the currently selected units of measure or the requested units of measure). This command is now handled by using the full scale query FPF (variableID) [engUnitsID], where (variableID) is the ID associated with the variable you would like to query the full scale of and [engUnitsID] is an optional input for querying the full scale in a particular unit of measure. If [engUnitsID] input is left blank, the command will return the full scale of the desired variable in the currently selected units of measure. Below is are examples of how this would look, both with and without the use of the optional [engUnitsID] input:

The returned string of the full scale query command will include the unit ID, the full scale as a floating point number, the engineering units ID associated with the queried variable, and the units of measure as a text string.

The following pages will serve as a guide for the data variable ID’s and the engineering unit ID’s associated with them.

| Variable Name | ID Value | Notes |
| --- | --- | --- |
| None | 1 | No statistic; leaves the location empty. |
| Absolute Pressure | 2 |  |
| Flow Temperature | 3 |  |
| Volumetric Flow | 4 | Volumetric flow referenced to flow conditions. |
| Mass Flow | 5 |  |
| Gauge Pressure | 6 |  |
| Differential Pressure | 7 |  |
| Total Volume | 8 | Totalized volume, referenced to flow conditions. |
| Total Mass | 9 | Totalized mass. |
| Totalizing Time | 10 | Amount of time that the totalizer has been running. |
| Remaining Batch Volume | 11 | Remaining volume, referenced to flow conditions, in the totalizer batch. |
| Remaining Batch Mass | 12 | Remaining mass in the totalizer batch. |
| Valve Drive | 13 |  |
| Remaining Measurement Time | 14 |  |
| Absolute External Pressure | 15 | Barometer reading, if a barometric pressure sensor is installed on the device. |
| External Temperature | 16 | Temperature probe reading. (N/A for standard devices) |
| External Volumetric Flow | 17 | Volumetric flow referenced to external conditions. (N/A for standard devices) |
| Total External Volume | 18 | Totalized volume, referenced to external conditions. (N/A for standard devices) |
| Remaining External Batch Volume | 19 | Remaining volume, referenced to external conditions, in the totalizer batch. (N/A for standard devices) |
| External Volumetric Flow Reference Temperature | 20 | Temperature probe reading if available, flow temperature reading otherwise. (N/A for standard devices) |
| External Volumetric Flow Temperature Source | 21 | Extrnl if using temperature probe for external volumetric temperature reference, Stream otherwise. (N/A for standard devices) |
| Flow Fraction | 22 | The fraction of current flow that is represented by the associated data source. (N/A for standard devices) |
| Total Fraction | 23 | The fraction of the totalizer value that is represented by the associated data source. (N/A for standard devices) |
| Sensor Fraction | 24 | The fraction of the current flow / volume (depends on how the sensor takes readings) that is represented by the associated data source. (N/A for standard devices) |
| Setpoint | 32 | The current setpoint reading. When specified in a location, this statistic will be replaced with the currently active setpoint statistic. |
| Absolute Pressure Setpoint | 34 |  |
| Volumetric Flow Setpoint | 36 | The setpoint for volumetric flow referenced to flow conditions. |
| Mass Flow Setpoint | 37 |  |
| Gauge Pressure Setpoint | 38 |  |
| Differential Pressure Setpoint | 39 |  |
| External Volumetric Flow Setpoint | 49 | The setpoint for volumetric flow, referenced to external conditions. (N/A for standard devices) |
| Average Volumetric Flow | 68 | Average volumetric flow, referenced to flow conditions, over the time of totalizing. |
| Average Mass Flow | 69 | Average mass flow over the time of totalizing. |
| Average External Volumetric Flow | 81 | Average volumetric flow, referenced to external flow conditions, over the time of totalizing. (N/A for standard devices) |
| Peak Volumetric Flow | 100 | Peak volumetric flow, referenced to flow conditions, during the time of totalizing. |
| Peak Mass Flow | 101 | Peak mass flow during the time of totalizing. |
| Peak External Volumetric Flow | 113 | Peak volumetric flow, referenced to external flow conditions, during the time of totalizing. (N/A for standard devices) |

Available pressure units:

Unit labels have an A appended for absolute pressure, a G appended for gauge pressure, or a D for differential pressure; unit scale factors remain unchanged across the different pressure types.

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| Pa | 2 | pascal |
| hPa | 3 | hectopascal |
| kPa | 4 | kilopascal |
| MPa | 5 | megapascal |
| mbar | 6 | millibar |
| bar | 7 | bar |
| g/cm2 | 8 | gram force per square centimeter |
| kg/cm | 9 | kilogram force per square centimeter |
| PSI | 10 | pound force per square inch |
| PSF | 11 | pound force per square foot |
| mTorr | 12 | millitorr |
| torr | 13 | torr |
| mmHg | 14 | millimeter of mercury at 0 C |
| inHg | 15 | inch of mercury at 0 C |
| mmH2O | 16 | millimeter of water at 4 C (NIST conventional) |
| mmH2O | 17 | millimeter of water at 60 C |
| cmH2O | 18 | centimeter of water at 4 C (NIST conventional) |
| cmH2O | 19 | centimeter of water at 60 C |
| inH2O | 20 | inch of water at 4 C (NIST conventional) |
| inH2O | 21 | inch of water at 60 C |
| atm | 22 | atmosphere (only available for absolute pressure) |
| m asl | 23 | meter above sea level (only available in /ALT builds and only for absolute pressure) |
| ft asl | 24 | foot above sea level (only available in /ALT builds and only for absolute pressure) |
| V | 61 | volt; no conversions are performed to or from other units |
| count | 62 | setpoint count, 0 – 64000 |
| % | 63 | percent of full scale |

Available volumetric flow units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| uL/m | 2 | microliter per minute |
| mL/s | 3 | milliliter per second |
| mL/m | 4 | milliliter per minute |
| mL/h | 5 | milliliter per hour |
| L/s | 6 | liter per second |
| LPM | 7 | liter per minute |
| L/h | 8 | liter per hour |
| US GPM | 9 | US gallon per minute |
| US GPH | 10 | US gallon per hour |
| CCS | 11 | cubic centimeter per second |
| CCM | 12 | cubic centimeter per minute |
| cm3/h | 13 | cubic centimeter per hour |
| m3/m | 14 | cubic meter per minute |
| m3/h | 15 | cubic meter per hour |
| m3/d | 16 | cubic meter per day |
| in3/m | 17 | cubic inch per minute |
| CFM | 18 | cubic foot per minute |
| CFH | 19 | cubic foot per hour |
| CFD | 21 | cubic foot per day |
| count | 62 | setpoint count, 0 – 64000 |
| % | 63 | percent of full scale |

Available total volume units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| uL | 2 | microliter |
| mL | 3 | milliliter |
| L | 4 | liter |
| US GAL | 5 | US gallon |
| cm3 | 6 | cubic centimeter |
| m3 | 7 | cubic meter |
| in3 | 8 | cubic inch |
| ft3 | 9 | cubic foot |
| uP | 61 | micropoise, a measure of viscosity; no conversions are performed to or from other units |

Available standard and normal flow units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| SuL/m | 2 | standard microliter per minute |
| SmL/s | 3 | standard milliliter per second |
| SmL/m | 4 | standard milliliter per minute |
| SmL/h | 5 | standard milliliter per hour |
| SL/s | 6 | standard liter per second |
| SLPM | 7 | standard liter per minute |
| SL/h | 8 | standard liter per hour |
| SCCS | 11 | standard cubic centimeter per second |
| SCCM | 12 | standard cubic centimeter per minute |
| Scm3/h | 13 | standard cubic centimeter per hour |
| Sm3/m | 14 | standard cubic meter per minute |
| Sm3/h | 15 | standard cubic meter per hour |
| Sm3/d | 16 | standard cubic meter per day |
| Sin3/m | 17 | standard cubic inch per minute |
| SCFM | 18 | standard cubic foot per minute |
| SCFH | 19 | standard cubic foot per hour |
| SCFD | 21 | standard cubic foot per day |
| kSCFM | 20 | 1000 standard cubic feet per minute |
| NuL/m | 32 | normal microliter per minute |
| NmL/s | 33 | normal milliliter per second |
| NmL/m | 34 | normal milliliter per minute |
| NmL/h | 35 | normal milliliter per hour |
| NL/s | 36 | normal liter per second |
| NLPM | 37 | normal liter per minute |
| NL/h | 38 | normal liter per hour |
| NCCS | 41 | normal cubic centimeter per minute |
| NCCM | 42 | normal cubic centimeter per minute |
| Ncm3/h | 43 | normal cubic centimeter per hour |
| Nm3/m | 44 | normal cubic meter per minute |
| Nm3/h | 45 | normal cubic meter per hour |
| Nm3/d | 46 | normal cubic meter per day |
| count | 62 | setpoint count, 0 – 64000 |
| % | 63 | percent of full scale |

Available total standard volume and normal volume units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| SuL | 2 | standard microliter |
| SmL | 3 | standard milliliter |
| SL | 4 | standard liter |
| Scm3 | 6 | standard cubic centimeter |
| Sm3 | 7 | standard cubic meter |
| Sin3 | 8 | standard cubic inch |
| Sft3 | 9 | standard cubic foot |
| kSft3 | 10 | 1000 standard cubic feet |
| NuL | 32 | normal microliter |
| NmL | 33 | normal milliliter |
| NL | 34 | normal liter |
| Ncm3 | 35 | normal cubic centimeter |
| Nm3 | 36 | normal cubic meter |

Available true mass flow units (these can be used for mass flow on gas devices or for volumetric flow on liquid devices calibrated with a true mass flow unit):

| Unit Label | Value | Notes |
| --- | --- | --- |
| mg/s | 64 | milligram per second |
| mg/m | 65 | milligram per minute |
| g/s | 66 | gram per second |
| g/m | 67 | gram per minute |
| g/h | 68 | gram per hour |
| kg/m | 69 | kilogram per minute |
| kg/h | 70 | kilogram per hour |
| oz/s | 71 | ounce per second |
| oz/m | 72 | ounce per minute |
| lb/m | 73 | pound per minute |
| lb/h | 74 | pound per hour |

Available total true mass units (these can be used for total mass on gas devices or for total volume on liquid devices calibrated with a total true mass unit):

| Unit Label | Value | Notes |
| --- | --- | --- |
| mg | 64 | milligram |
| g | 65 | gram |
| kg | 66 | kilogram |
| oz | 67 | ounce |
| lb | 68 | pound |

Available temperature units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| °C | 2 | degree Celsius |
| °F | 3 | degree Farenheit |
| K | 4 | Kelvin |
| °Ra | 5 | degree Rankine |

Available valve drive units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
| count | 2 | +/- 65536 at full drive |
| % | 63 | percent of full scale drive |

Available time units:

| Unit Label | Value | Notes |
| --- | --- | --- |
|  | 0 | Unit not specified; use default values. |
| --- | 1 | Unknown unit; no conversions are performed to other units. If calibrated with different units, the value is displayed in those units. |
|  | 2 | Value displayed as hours:minutes:seconds |
| ms | 3 | millisecond |
| s | 4 | second |
| m | 5 | minute |
| hour | 6 | hour |
| day | 7 | day |